焦度计项目 AI 开发总规范（AI 看版）

先确认板块和阶段，再按统一接口开发

AI 开工前必须先问：1. 你负责哪个板块？M1 输入与配置 / M2 图像识别 / M3 标定与计算 / M4 本地系统与展示。 2. 当前处于哪个阶段？第一阶段模块成品 / 第二阶段合并修 bug / 第三阶段优化创新。 3. 你要我先做什么？代码、接口、测试样例、README、调试、报告，还是其他？ 4. 你现在有哪些本地文件？标定图、测量图、配置文件、mock 数据分别在哪里？

# 四大模块边界

| 模块 | 允许做 | 必须输出 | 禁止做 |
| --- | --- | --- | --- |
| M1 输入与配置 | 读取图片路径、配置参数、样本列表、任务编号 | 标准输入数据包 input_package.json | 禁止做图像识别、禁止计算 S/C/A |
| M2 图像识别 | ROI、滤波、增强、二值化、连通域、质心 | spots_calib.json 和 spots_meas.json | 禁止计算镜片度数 |
| M3 标定与计算 | 坐标系、光斑位移、镜片类型、S/C/A、质量判断 | result.json | 禁止读取原始图片重新识别 |
| M4 本地系统与展示 | 主流程、界面/命令行、日志、结果展示、导出、异常提示 | 可操作本地系统 | 禁止重写 M2/M3 核心算法 |

# 统一目录结构

focimeter_system/

config/default_config.json

data/samples/calibration/

data/samples/measurement/

data/mock/

modules/input_config/

modules/image_recognition/

modules/calibration_calculation/

modules/local_system/

outputs/images/ outputs/results/ outputs/logs/ outputs/reports/

# 统一输入数据包

{

"task_id": "sample_001",

"calibration_image": "data/samples/calibration/calib_001.png",

"measurement_image": "data/samples/measurement/meas_001.png",

"config_path": "config/default_config.json",

"run_mode": "local_image"

}

# 统一配置文件

{

"camera": {"pixel_size_um": 4.0, "image_width": null, "image_height": null},

"optical": {"distance_m": 0.03, "hartmann_spacing_mm": null},

"image_processing": {"roi_width_ratio": 0.9, "roi_height_ratio": 0.9, "median_kernel": 3, "tophat_kernel": 30, "otsu_a": 0.4, "otsu_b": 0.7, "max_depth": 2},

"recognition": {"expected_spot_count": 5, "min_confidence": 0.7},

"calculation": {"pixel_threshold": 1.0, "angle_unit": "degree"}

}

# 统一光斑识别输出

{

"task_id": "sample_001", "status": "ok",

"image_type": "calibration", "coordinate_type": "image_pixel",

"spots": [{"spot_id": 0, "role": "center", "x": 512.34, "y": 384.21, "confidence": 0.96}],

"quality": {"expected_count": 5, "detected_count": 5, "is_usable": true},

"error": null

}

统一 role 值：center、y_positive、x_positive、left_or_negative、other、unknown。

# 统一计算输出

{

"task_id": "sample_001", "status": "ok", "lens_type": "spherical",

"result": {"S": -2.50, "C": 0.00, "A": null, "unit": "D"},

"quality": {"is_usable": true, "confidence": 0.91, "warnings": []},

"intermediate": {"coordinate_system_valid": true, "shift_unit": "pixel"},

"error": null

}

# 统一错误格式

{

"status": "error",

"error": {"code": "SPOT_COUNT_MISMATCH", "message": "Expected 5 spots but detected 3.", "module": "image_recognition", "recoverable": true}

}

# 统一单位

| 物理量 | 统一单位 | 字段写法 |
| --- | --- | --- |
| 图像坐标 | pixel | x、y，且 coordinate_type=image_pixel |
| 像元尺寸 | um | pixel_size_um |
| 物理距离 | m | distance_m |
| 光阑间距 | mm | hartmann_spacing_mm |
| 角度 | degree | A、angle_unit=degree |
| 屈光度 | D | S、C、unit=D |
| 置信度 | 0 到 1 | confidence=0.91，不用百分数 |

# 统一坐标系

图像坐标：左上角为原点，X 向右为正，Y 向下为正，单位 pixel。

标定坐标：中心光斑为原点，y_positive 定义 Y 轴正方向，x_positive 定义 X 轴正方向。M3 计算前必须说明坐标是否已经从图像坐标转换到标定坐标。

# AI 开发硬性要求

先问成员负责哪个板块，再开始。

先问当前阶段，再开始。

需要其他模块时，先用 mock 数据。

不能擅自改字段名、单位、坐标系、错误码。

不确定的硬件参数必须标注 TODO_CONFIRM。

每次输出都要包含运行方法、输入样例、输出样例、错误样例。